List of Tables

Table 1. The composition of the NiTi alloy, AISI 316L austenitic stainless steel and X2CrNiMoN22-5-3 duplex stainless steel, wt.%

Table 2. The electrochemical parameters obtained from potentiodynamic polarization measurements for NiTi alloy, AISI 316L austenitic stainless steel and X2CrNiMoN22-5-3 duplex stainless steel in 0.9 % sodium chloride solution

Table 3. Parameters of equivalent circuit for NiTi alloy, X2CrNiMoN22-5-3 duplex stainless steel and AISI 316L austenitic stainless steel in 0.9% sodium chloride solution after EIS measurements

Table 1. The composition of the NiTi alloy, AISI 316L austenitic stainless steel and X2CrNiMoN22-5-3 duplex stainless steel, wt.%

| Material | Chemical Composition (wt.%) |
| --- | --- |
| NiTi alloy | 49.16% Ti, 49.75% Ni |
| AISI 316L austenitic stainless steel | 16.64 % Cr, 10.68 % Ni, 2.53 % Mo, 0.47 % Si, 0.002% C, 0.024% Nb, 0.009% S, 0.0045 % P, 69.60 % Fe |
| X2CrNiMoN22-5-3 duplex stainless steel | 22.21% Cr, 5.47% Ni, 3.14% Mo, 1.52% Mn, 0.33% Si, 0.26 % Cu, 0.012 % C, 0.035 % V, 0.012 % Nb, 67.00 %Fe |

Table 2. The electrochemical parameters obtained from potentiodynamic polarization measurements for NiTi alloy, AISI 316L austenitic stainless steel and X2CrNiMoN22-5-3 duplex stainless steel in 0.9 % sodium chloride solution

| Material | icorr / A cm-2 | Ecorr / V | bc / V dec-1 | ba / V dec-1 | Eb / V | vcorr / mm a-1 |
| --- | --- | --- | --- | --- | --- | --- |
| NiTi alloy | 0.42 | -0.224 | -0.2508 | 0.2934 | 0.470 | 0.00364 |
| AISI 316L austenitic stainless steel | 1.05 | -0.234 | -0.1309 | 0.2336 | 0.245 | 0.0109 |
| X2CrNiMoN22-5-3 duplex stainless steel | 0.68 | -0.240 | -0.127 | 0.189 | - | 0.00719 |

Table 3. Parameters of equivalent circuit for NiTi alloy, X2CrNiMoN22-5-3 duplex stainless steel and AISI 316L austenitic stainless steel in 0.9% sodium chloride solution after EIS measurements

| Material | Time, hours | Qp 106 (-1sncm-2) | n1 | Rp (kcm2) | Qb 106 (-1sn cm-2) | nb | Rb (k cm2) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| NiTi alloy | 1 | 71.48 | 0.89 | 82.57 | 50.52 | 0.91 | 414.83 |
|  | 3 | 65.40 | 0.88 | 91.48 | 44.68 | 0.91 | 552.31 |
|  | 6 | 62.35 | 0.89 | 98.48 | 40.82 | 0.92 | 646.47 |
|  | 12 | 60.65 | 0.88 | 108.88 | 39.40 | 0.92 | 728.20 |
|  | 24 | 56.78 | 0.89 | 120.00 | 37.56 | 0.92 | 704.12 |
| X2CrNiMoN22-5-3 duplex SS [a] | 1 | 93.98 | 0.88 | 50.68 | 78.74 | 0.90 | 272.90 |
|  | 3 | 76.22 | 0.88 | 61.29 | 56.81 | 0.91 | 349.35 |
|  | 6 | 67.39 | 0.88 | 65.44 | 58.62 | 0.91 | 403.22 |
|  | 12 | 67.53 | 0.87 | 76.59 | 53.53 | 0.91 | 442.44 |
|  | 24 | 64.38 | 0.89 | 97.28 | 52.06 | 0.89 | 476.79 |
| AISI 316L austenitic SS[a] | 1 | 130.78 | 0.86 | 40.49 | 110.87 | 0.89 | 55.65 |
|  | 3 | 104.31 | 0.86 | 48.80 | 95.51 | 0.90 | 102.51 |
|  | 6 | 96.34 | 0.87 | 54.44 | 91.94 | 0.91 | 117.65 |
|  | 12 | 94.54 | 0.87 | 65.74 | 85.32 | 0.90 | 147.65 |
|  | 24 | 88.39 | 0.87 | 80.41 | 80.48 | 0.90 | 185.07 |

[a] SS – stainless steel